## Slide 1: Instrument intercomparison in the stray radiation field around the CERN proton synchrotron
E. Aza, M. Caresana, C. Cassell, V. Colombo, S. Damjanovic, S. Gilardoni,
G.P. Manessi, M. Pangallo, D. Perrin, M. Silari
3rd June 2013, Aix-en-Provence
12th NEUtron and ion DOSimetry symposium

## Slide 2
THE PS COMPLEX
TO SPS
BEAM
INJECTION
BEAM
EXTRACTION
E
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
1/11
100 combined-function magnets interleaved by 100 straight sections (SS)
Major losses due to beam injection and extraction
Beam accelerated from 1.4 GeV to 26 GeV

## Slide 3
THE MEASUREMENT CAMPAIGN
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
2/11
Aim of the measurements: evaluate the difference in the responses of 6 active neutron detectors in different stray field conditions.
Focus: difference in the responses due to dead time losses.
4 MEASURING
POSITIONS
Non-pulsed losses (200-300 μs), smoothed by 6-10 meters of ground shield
Pulsed losses, low intensity (quite far from the
extraction septum)
Pulsed, intense losses (access tunnel to the extraction septum, limited shielding)
Several measurements carried out in 4 the positions
Systematic intercomparison at the access tunnel of SS16

## Slide 4
THE INTERCOMPARISON
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
3/11
1% lost beam in the area: losses due to particles scattered by the electrostatic septum used to slice the beam
Beam extracted in 5 turns: each turn lasts 2.2 μs => length of the losses: 11 μs
MEASURING POINT
6 detectors installed and rotated in reference positions at the beginning of the access tunnel
Beam extracted with Continuous Transfer (CT) technique

## Slide 5
THE DETECTORS (I)
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
4/11
3 classes of instruments employed:
I) Conventional and extended range rem counters
Used for RP routine measurements
II) Ionization chambers
Employed in RAMSES (CERN RAdiation Monitoring System for the Environment and Safety)
III) Prototype extended range rem counter (LUPIN)
Thermo BIOREM (BF3)
Conventional rem counter
Good sensitivity up to 20 MeV
Thermo Wendi-2 (3He)
Extended range rem counter
Good sensitivity up to 5 GeV
CERN LINUS (3He)
The original extended range rem counter

## Slide 6
THE DETECTORS (II)
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
5/11
Pressurized (20 atm) Centronic ionization chambers (Ar and H)
Good sensitivity up to 20 MeV
LUPIN prototype (BF3)
Extended range rem counter
Specifically conceived to work in pulsed neutron fields
It uses the integration of the total electric charge as a measure of the number of neutron interactions in the gas
Chart: Time [μs]

## Slide 7
THE STRAY FIELD
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
6/11
Example of signal acquired with the LUPIN in one of the reference positions
THERMALIZATION &
DIFFUSION TIME
250 INTERACTIONS IN 300 μs
NEUTRONS SCATTERED ALONG THE TUNNEL WALLS

## Slide 8
THE NEUTRON SPECTRUM
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
7/11
Expected neutron spectrum at the beginning of the access tunnel (FLUKA simulations)
Expected neutron H*(10) rate = 95% of the total H*(10) rate
Max neutron energy < 20 MeV => Same response for conventional and extended range RC

## Slide 9
THE COMPARISON PROCESS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
8/11
Detectors turned in the 6 reference positions
(high H*(10) rate gradient in the area)
Comparison of the integrated H*(10) (30 minutes) in each positions
EXPECTED H*(10) [μSv/lost primary]
Normalization of the H*(10) to:
Integrated proton fluence in the PS
H*(10) integrated by a fixed monitoring station installed in the area

## Slide 10
THE RESULTS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
9/11
REM COUNTERS
DEAD TIME ≈ 2 μs
REM COUNTERS
DEAD TIME ≈ 1 μs
ION. CHAMBERS
& LUPIN
(NO DEAD TIME)
Data normalized to the H*(10) integrated by the fixed station
Non-constant losses during measurements (different beam cycles)
0
FLUKA
LOWER H*(10)
(CLOSE TO WALL)
Notes: Uncertainties, not shown for clarity: 5% positioning, ≈1% statistical.

## Slide 11
REMARKS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
10/11
1) Ionization chambers & LUPIN (no dead time)
Results compatible within the range of uncertainty and with FLUKA simulations.
2) BIOREM (dead time ≈ 1 μs)
30% underestimation if compared to the average results of group 1 and FLUKA
1) Wendi-2 & LINUS (dead time ≈ 2 μs)
60% underestimation if compared to the average results of group 1 and FLUKA
From the dead time correction formula:
n = expected counting rate ≈ 8.3∙105 s-1 (from LUPIN signal)
m = measured counting rate
τ = dead time ≈ 1 or 2∙10-6 s
Expected underestimation due to the dead time effects:
45% for BIOREM; 65% for Wendi-2 & LINUS

## Slide 12
CONCLUSIONS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
11/11
LUPIN and IONIZATION CHAMBERS (H/Ar)
No underestimation of the expected H*(10)
Can be used in every working conditions, even with strongly pulsed fields.
Employed at CERN in all critical positions (pulsed, high energy losses).
LUPIN has the advantage of being portable (rem counter) and (unlike the ionization chambers) does not need the setting of a different calibration factor in each measuring position as a function of the neutron spectrum.
CONVENTIONAL REM COUNTERS (BIOREM, LINUS, Wendi-2)
Huge underestimations due to dead time
Real underestimation slightly lower than what calculated with the DT correction formula
(stray field partly scattered)
Used at CERN for environmental measurements or outside accelerator shieldings where radiation field is not strongly pulsed.

## Slide 13: Instrument intercomparison in the stray radiation field around the CERN proton synchrotron
THANK YOU FOR YOUR ATTENTION
E. Aza, M. Caresana, C. Cassell, V. Colombo, S. Damjanovic,
S. Gilardoni, G.P. Manessi, M. Pangallo, D. Perrin, M. Silari

## Slide 14
DEAD TIME CALCULATIONS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
Backup 1
To derive the dead time for each detector the “two-source method” [Knoll] was applied to the data acquired in a dedicated experiment.
Method based on observing the counting rate from two sources individually and in combination. Applied by taking three different values of expected and measured counting rates, i.e. EXP1, EXP2, EXP12 and MEAS1, MEAS2, MEAS12, where EXP12 = EXP1 + EXP2, and by evaluating them individually and in combination.
The dead time could be calculated from the discrepancy between the measured rate MEAS12 and the sum of MEAS1 + MEAS2.
(1.02 ± 0.10) µs for the BIOREM
(very close to the TTL pulse width declared by the manufacturer, 1.2 µs [Thermo])
(1.74 ± 0.17) µs for the Wendi-2

## Slide 15
RESPONSE FUNCTIONS
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
Backup 2
Chart: Neutron energy [MeV]
Chart: Neutron energy [MeV]
Chart: Neutron energy [MeV]
H20
Chart: Energy [MeV]
| 1.0000000000000048E-8 | 2.0000000000000209E-8 | 3.0000000000000269E-8 | 4.0000000000000272E-8 | 5.0000000000000342E-8 | 6.0000000000000511E-8 | 7.0000000000000521E-8 | 8.0000000000000571E-8 | 9.00000000000007E-8 | 1.0000000000000043E-7 | 1.0000000000000033E-6 | 1.0000000000000031E-5 | 1.0000000000000025E-4 | 1.0000000000000024E-3 | 4.0000000000000114E-3 | 5.0000000000000114E-3 | 6.0000000000000114E-3 | 7.0000000000000114E-3 | 8.0000000000000227E-3 | 1.0000000000000005E-2 | 2.0000000000000011E-2 | 3.0000000000000002E-2 | 0.05 | 7.0000000000000021E-2 | 0.1 | 0.2 | 0.30000000000000032 | 0.4 | 0.5 | 0.60000000000000164 | 0.70000000000000095 | 0.8 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 100 | 200
Biorem Response Function | 1.1 | 1.1 | 1.15 | 1.15 | 1.175 | 1.2 | 1.22 | 1.23 | 1.24 | 1.4 | 2.1 | 3.8 | 5.3 | 6.8 | 7.7 | 7.4 | 7.2 | 7 | 6.7 | 6.4 | 5 | 3 | 2 | 1.6 | 1.1 | 0.9 | 0.76 | 0.76 | 0.8 | 0.9 | 0.99 | 1 | 1.01 | 1.05 | 1.01 | 1 | 0.9 | 0.8 | 0.75 | 0.7 | 0.54 | 0.3 | 0.24 | 0.235 | 0.201 | 0.2 | 0.18 | 0.175 | 0.16 | 0.145
A20
Wendi-2
BIOREM
Chart: Neutron energy [GeV]
LINUS
Chart: Energy [GeV]
| 9.9999999999999986E-10 | 9.9999999999999995E-7 | 1E-4 | 1E-3 | 5.0000000000000001E-3 | 0.01 | 0.02 | 0.05 | 0.1 | 0.3 | 1
 | 3.2e-07 | 7.7e-07 | 1.9e-06 | 1.06e-05 | 1.4e-05 | 1.2e-05 | 7.9e-06 | 6.1e-06 | 5.6e-06 | 6.1e-06 | 1.09e-05

LUPIN

## Slide 16
NORMALIZATION TO PS FLUENCE
G.P. Manessi – Instrument intercomparison in the stray radiation field around the CERN PS
Backup 3
Data normalized to integrated proton fluence in the PS